BRCA1-like MLPA classifier

This online tool helps you to classify a tumor sample as BRCA1-like or Sporadic-like, when the P376-BRCA1ness probe mix of MRC-Holland is used1.

The BRCA1-like signature is derived from BRCA1-mutated tumors. The BRCA1-like aCGH pattern may also be present in sporadic breast cancers, which frequently show evidence of impaired BRCA1 function due to other causes than mutation, for example, methylation. We and others have shown that sporadic basal-like breast cancers resemble BRCA1 mutated cancers in many different ways. This concept of "BRCAness" in sporadic cancers is under study in many centres as it may have implications for treatment selection.

- Input data

Input should be the output of Genemapper (or any other fragment analysis software program) with the samples in rows and the height and peak areas as columns.

Explanation

This input file should contain between 3 and 10 Reference samples (reference samples should be named R1, R2, R3, etc.) and peak heights (for QC) and peak areas (assay probes). The peaks should be named Peak.Height 1…till Peak. Area 54. Be aware that the probes with length 255, 445 and 500 are of insufficient quality and should be removed in the genemapper output. The first 9 probes (or Peak Heights) are for quality control. Starting from probe 10 (Peak Area 10) are the assay probes. Check the example Raw input file carefully and be sure that your input file has exactly the same layout.

The tool first checks the quality of the values by two quality checks (see also the general MLPA guidelines):

- QC1: If the peak heights of the Q-fragments (64,70,76,82 nt) (Peak.Heights 1,2,3,and 4) are all higher than 1/3 of the 92nt (Peak.Height 6) fragment, than the quantity of the input DNA is not sufficient, or the ligation step failed. Repeat the experiment for this sample with sufficient DNA.

- QC2: If one of the D-fragments (88, 96 nt) is much lower (<40%) than the 92nt fragment, the denaturation of the DNA failed. Repeat the experiment for this sample.

- PAM input

This sheet contains the normalized MLPA data, either from step 1 (Input data) or your own normalized data.

Explanation

This sheet contains the normalized MLPA data. In the first normalization step the probes are divided by the sum of all probes for that specific sample. In the second normalization step each probe value per sample is divided by the mean value of the reference probes. After normalization, all 34 BRCA1-like classification probes are put together into this “Pam input“ file. Alternatively, if you have performed the normalization yourself, you can upload your normalized data here. In this case, you should check that your PAM input file has exactly the same lay-out as the example Normalized Data.

- Classification

This sheet contains the BRCA1-like classification result.

Explanation

This sheet contains the BRCA1-like classification. Each sample gets a BRCA1-like and a Sporadic-like score, we call a sample with a BRCA1-like score of 0.5 or higher BRCA1-like. This sheet also contains the results of the quality control.

Reference

- Lips EH, Laddach N, Savola SP, Vollebergh MA, Oonk AM, Imholz AL, Wessels LF, Wesseling J, Nederlof PM, Rodenhuis S. Quantitative copy number analysis by Multiplex Ligation-dependent Probe Amplification (MLPA) of BRCA1-associated breast cancer regions identifies BRCAness. Breast Cancer Res. 2011;13:R107.